## FMEA
|  | Company Name: |  |  |  |  | Engineering Location: | [5 empty] | Customer Name:
|  | Model Year(s) / Program(s): |  |  |  |  | Subject: | [5 empty] | FMEA Start Date:
|  | FMEA Revision Date: |  |  |  |  | Cross-Functional Team: | [5 empty] | FMEA DOC/ID Number:
|  | Design/Process Owner: |  |  |  |  | Confidentiality Level: | [5 empty] | FMEA Type: | ISO 14971 complient FMEA
Focus Element / Process / System:
|  | FF (Step 1) | Failure Analysis (Step 4) |  |  |  | Risk Analysis (Step 5) |  |  | Optimization (Step 6) | [9 empty] | Propability Analysis
ID | UID | Focus Function | Failure Effect (FE) | Severity (S) of FE | Failure Mode (FM) of the Focus Element | Failure Cause (FC) Chain | Current Prevention Control (PC) of FC | Occurrence (O) of FC | Risk Level (RL) | Preventiv Action | Detection Action | Responsible Person's Name | Target Completion Date | Status | Action Taken with Pointer to Evidence | Completion Date | Severity (S) | Occurrence (O) | Risk Level (RL) | P1 | P2 | O | AL
[9 empty] | 0 | [9 empty] | 0

## Severity
Severity (S) of Harm Assessment List
S | Effect | Impact | Corporate Examples
5 | Catastrophic / Fatal | Description: Results in death of one or more patients, users, or other persons. Examples: Electrocution, massive hemorrhage, organ failure, device explosion, inability of a life-sustaining device to perform its function leading directly to death. | Fill in own examples
4 | Critical / Permanent Impairment / Life-Threatening | Results in permanent impairment of a body structure or function, or life-threatening injury/illness requiring immediate medical or surgical intervention to prevent death. Examples: Irreversible brain damage, loss of a limb, permanent blindness, severe allergic reaction requiring emergency care, cardiac arrest from which patient is resuscitated but with permanent damage. | Fill in own examples
3 | Serious / Major Injury / Significant Intervention | Description: Results in injury or impairment requiring medical or surgical intervention to prevent permanent impairment (but not life-threatening), or significant degradation in health requiring extended hospitalization. Examples: Fractures requiring surgery, severe burns requiring skin grafts, infection requiring IV antibiotics and hospitalization, significant blood loss requiring transfusion, major device malfunction requiring patient re-operation. | Fill in own examples
2 | Minor / Temporary Injury / Minor Intervention | Results in temporary injury or impairment not requiring medical or surgical intervention, or requiring minimal, non-invasive intervention (e.g., first aid, over-the-counter medication, short observation). Examples: Minor skin irritation, bruising, temporary headache, transient pain, short-term discomfort, minor incision site infection treated with oral antibiotics. | Fill in own examples
1 | Negligible / Inconvenience / Temporary Discomfort | Results in inconvenience or temporary discomfort with no lasting effects or need for medical intervention. Examples: Slight discomfort, minor bruising that resolves quickly, temporary cosmetic issue, device alarm sounding unnecessarily but no harm ensues, delay in a non-critical diagnostic procedure. | Fill in own examples
While FMEA severity focuses on the "effect of the failure mode", ISO 14971 risk management focuses on the "harm" that could result from a hazardous situation. It's often expressed as the "measure of the possible consequences of a hazard". This distinction is crucial: a design failure might lead to a hazardous situation, but the harm is the physical injury or damage to the health of people, or damage to property or the environment resulting from that situation. ISO 14971 itself does not prescribe a specific severity scale (e.g., 1-5 or 1-10) because the appropriate scale can vary depending on the device, its intended use, and the patient population. However, it does require the manufacturer to define and document their own severity scale within their risk management plan.
Attention: The scale shown here is only an example. Each manufacturer of medical devices must define its own scale to suit its medical device. An injection cannula that has been manufactured and used millions of times is not the same as an innovative system that has just come onto the market, such as e.g. an Impella cardiac catheter that can save lives that were previously beyond saving.
Key Aspects for ISO 14971 Severity Definition:
Focus on Harm: Explicitly links the severity to the harm to people (patient, user, third party) or the environment, not just the device failure.
Specific and Objective Definitions: The standard emphasizes that severity definitions should be clear, unambiguous, and, where possible, objective, to ensure consistent application across different risk analyses within the organization. Avoid vague terms like "high" or "low" without further description.
Intended Use and Patient Population: The severity of harm can vary significantly based on the intended use of the device and the characteristics of the patient population (e.g., an issue might be minor for a healthy adult but catastrophic for an infant or an immunocompromised patient). This should be considered when defining or applying the scale.
Reasonably Foreseeable Misuse: ISO 14971 also requires consideration of harm from reasonably foreseeable misuse, not just normal use.
Worst-Case Harm: When assessing a hazardous situation, the severity should typically reflect the worst reasonably foreseeable harm that could occur.

## Occurrence
Probability (P) of Occurrence of Harm
S | Effect | Impact | Corporate Examples
5 | Frequent / Probable | The hazardous situation leading to harm is expected to occur frequently during the lifetime of a device or across a large number of devices/users. Likelihood: >1 in 10 to 1 in 2 (>10% to 50%) Examples: Well-documented, recurring issues in similar devices; design characteristics that consistently lead to user error and harm. | Fill in own examples
4 | Occasional / Reasonably Probable | The hazardous situation leading to harm is expected to occur occasionally during the lifetime of a device or across the device population. Likelihood: 1 in 100 to 1 in 10 (1% to 10%) Examples: Some reported instances in clinical trials or post-market surveillance of similar devices; known failure mode that occurs but not consistently. | Fill in own examples
3 | Remote / Unlikely | The hazardous situation leading to harm is unlikely to occur during the lifetime of a device or across the device population, but is not impossible. Likelihood: 1 in 1,000 to 1 in 100 (0.1% to 1%) Examples: Rare occurrences observed in extensive testing or very large field populations; theoretical failure modes with minimal supporting evidence. | Fill in own examples
2 | Improbable / Very Unlikely | The hazardous situation leading to harm is very unlikely to occur during the lifetime of a device or across the device population. Requires multiple, independent failures to coincide. Likelihood: 1 in 10,000 to 1 in 1,000 (0.01% to 0.1%) Examples: Highly robust design or process controls are in place; no known failures in extensive field history. | Fill in own examples
1 | Extremely Improbable / Negligible | The hazardous situation leading to harm is extremely unlikely to occur; considered practically impossible under normal or foreseeable conditions. Likelihood: <1 in 10,000 (<0.01%) Examples: Failure prevented by redundant systems; failure mode physically impossible due to design; extensive, mature use of Poka-Yoke or inherent safety features that effectively eliminate the possibility. | Fill in own examples
Just like Severity, ISO 14971:2019 does not prescribe a specific scale for the probability of occurrence. Instead, it mandates that the manufacturer define and justify their own scale within their risk management process. In the context of ISO 14971, "Occurrence" typically refers to the probability of a hazardous situation occurring, or, more precisely, the probability of the harm occurring. This is often a function of: 1. The probability of the initiating event or failure. 2. The probability of exposure to the resulting hazardous situation. 3. The probability of harm occurring given that exposure.
Attention: The scale shown here is only an example. Each manufacturer of medical devices must define its own scale to suit its medical device. An injection cannula that has been manufactured and used millions of times is not the same as a complicated system where only a few 10 thousand units are produced per year BUT has the potential to save lives that were previously beyond saving - such as e.g. an Impella cardiac catheter.
Key Aspects for ISO 14971 Severity Definition:
Life Cycle Perspective: Consider the probability of occurrence throughout the entire lifecycle of the device (design, manufacturing, transport, storage, installation, use, maintenance, disposal).
Existing Risk Control Measures: The probability rating should reflect the likelihood after considering the effectiveness of any already implemented risk control measures (e.g., existing design features, manufacturing controls, warnings, training). This is crucial for risk estimation.
Data Sources (Evidence-Based):
Clinical Data: Data from clinical trials, post-market surveillance of similar devices, literature.
Historical Data: Incident reports, warranty claims, complaint data for similar devices or components.
Testing Data: Results from verification and validation testing (e.g., accelerated life testing, stress testing, usability testing).
Simulation/Modeling: Reliability predictions, fault tree analysis (FTA), event tree analysis (ETA).
Expert Opinion: Justified expert judgment, particularly for novel designs or technologies where historical data is scarce.
Device Complexity and Criticality: More complex devices or those with highly critical functions may have higher inherent probabilities of certain failures.
User Interface and Human Factors: The design of the user interface and the potential for human error (usability issues) significantly influence the probability of a hazardous situation.
Environment of Use: The intended environment (e.g., home vs. hospital, sterile vs. non-sterile, extreme temperatures) can impact failure rates.
Single Point of Failure vs. Multiple Failures: Differentiate between hazardous situations resulting from a single failure versus those requiring a sequence of independent failures. Sequences of independent failures generally have a lower probability.
Traceability: Ensure that the assigned probability ratings are justified by traceable evidence and documented in the risk management file.

## Classification
Risk Level (RL)
Probability | P | Risk Assessment Matrix | [5 empty] | Probability | (P) | Risk Assessment Matrix
Frequent | 5 | A | U | U | U | U |  | Frequent | 5 | L | H | U | U | U
Occasional | 4 | A | A | U | U | U |  | Occasional | 4 | L | M | H | U | U
Remote | 3 | A | A | A | U | U |  | Remote | 3 | L | L | M | H | U
Improbable | 2 | A | A | A | A | U |  | Improbable | 2 | L | L | L | M | H
Extremely Improbalbe | 1 | A | A | A | A | A |  | Extremely Improbalbe | 1 | L | L | L | L | M
|  | 1 | 2 | 3 | 4 | 5 | S |  |  | 1 | 2 | 3 | 4 | 5 | S
|  | Negligible | Minor | Serious | Critical | Catastrophic | Severity |  |  | Negligible | Minor | Serious | Critical | Catastrophic | Severity
This is the generally accepted approach | [7 empty] | This is an approach that allows a finer division. However, be prepared for audits to question “green” areas - they do NOT mean that they should not be reduced further if at all possible - they just have a lower risk level.
|  |  | U | Unacceptable risk | [6 empty] | Unacceptable (U) - Red Zone | Action Required: These risks are generally not acceptable and require immediate risk reduction measures. Significant effort must be expended to reduce either the severity or the probability (or both) until the risk falls into an acceptable zone. If the risk cannot be reduced, the product may not be viable. The risk versus the benefit is unacceptable
|  |  | A | (Potentially) Acceptable risk - as low as reasonable possible | [6 empty] | High (H) - Orange Zone | Action Required: These risks are undesirable and require substantial risk reduction measures. All reasonably possible efforts should be made to reduce these risks. Requires management review and approval for acceptance if reduction is not feasible below this level. The risk versus the benefit is in general not acceptable.
[11 empty] | Medium (M) - Yellow Zone | Action Required: These risks are tolerable but should be monitored and considered for reduction if reasonably possible. Risk reduction measures are encouraged but may not be mandatory. Requires clear rationale for acceptance. The risk versus the benefit is acceptable.
[11 empty] | Low (L) - Green Zone | Action Required: These risks are considered acceptable without further risk reduction measures, provided they are effectively controlled and monitored. The risk versus the benefit is acceptable.
Key Considerations for Using a Risk Assessment Matrix:
Define Your Own Scale and Matrix: This example is common, but your organization must define its own specific severity and probability scales, the matrix structure, and the acceptance criteria. These definitions should be justified and documented in your risk management plan.
Risk Acceptance Criteria: A core requirement of ISO 14971 is to establish risk acceptance criteria before the risk analysis. This matrix is the visual representation of those criteria. The acceptable and unacceptable zones define what risks your organization is willing to tolerate.
Risk Estimation vs. Risk Evaluation:
You use the matrix to estimate risk based on initial (pre-control) or current (post-control) S and P values.
You then evaluate risk against your defined acceptance criteria to determine if further risk control is necessary.
As Low As Reasonably Practicable (ALARP): Even if a risk falls into a "tolerable" or "acceptable" zone, ISO 14971 expects manufacturers to reduce risks "as far as reasonably practicable" (ALARP principle), especially for risks related to safety.
Attention! Since the 2012 edition of EN ISO 14971, there are no longer any generally acceptable risks. The division into acceptable, unacceptable and ALARP risks is therefore obsolete. Instead, risks must be “as low as reasonably possible”. Be careful when using green ;)
Documentation: The entire risk management process, including the matrix and its interpretation, must be thoroughly documented in the device's Risk Management File.
Review and Update: The matrix and the risk acceptance criteria should be periodically reviewed and updated based on new information, standards, or organizational learning.
Residual Risks
ISO 14971 requires manufacturers to assess the acceptability of each individual residual risk. This assessment must be based on the criteria defined in the risk management plan.
Manufacturers are also obliged to ensure that the benefits of the product outweigh the overall residual risks.
Manufacturers must also decide which residual risks to disclose, e.g. in instructions for use.
Finally, the standard requires that residual risks and their acceptance be continuously assessed for as long as the products are on the market.

## ClassLogic
P | S | RL
1 | 1 | A
2 | 1 | A
3 | 1 | A
4 | 1 | A
5 | 1 | A
1 | 2 | A
2 | 2 | A
3 | 2 | A
4 | 2 | A
5 | 2 | U
1 | 3 | A
2 | 3 | A
3 | 3 | A
4 | 3 | U
5 | 3 | U
1 | 4 | A
2 | 4 | A
3 | 4 | U
4 | 4 | U
5 | 4 | U
1 | 5 | A
2 | 5 | U
3 | 5 | U
4 | 5 | U
5 | 5 | U